## Alex-Mn_As_data
Sample | [6 empty] | 27 Al [ He ] |  | 55 Mn [ He ] |  | 56 Fe [ He ] |  | 57 Fe [ He ] |  | 75 As [ He ] |  | 45 Sc ( ISTD ) [ He ] |  | 103 Rh ( ISTD ) [ He ] |  | 193 Ir ( ISTD ) [ He ]
| Rjct | Data File | Acq. Date-Time | Type | Level | Sample Name | Conc. [ ppb ] | Conc. RSD | Conc. [ ppb ] | Conc. RSD | Conc. [ ppb ] | Conc. RSD | Conc. [ ppb ] | Conc. RSD | Conc. [ ppb ] | Conc. RSD | Conc. | Conc. RSD | Conc. | Conc. RSD | Conc. | Conc. RSD
| FALSE | 001SMPL.d | 2023-11-16 11:31 | Sample |  | blk
56 Fe [ He ] : CPS RSD value = 5.02 is over the allowed maximum = 5.00 | FALSE | 002SMPL.d | 2023-11-16 11:33 | Sample |  | blk
| FALSE | 003SMPL.d | 2023-11-16 11:35 | Sample |  | blk
| FALSE | 004CALB.d | 2023-11-16 11:36 | CalBlk | 1 | 0 ppb | 0 | N/A | 0 | N/A | 0 | N/A | 0 | N/A | 0 | N/A
| FALSE | 005CALS.d | 2023-11-16 11:38 | CalStd | 2 | 0.5 ppb m16 + Al | <0.000 | N/A | 0.48 | 6 | <0.000 | N/A | <0.000 | N/A | 0.498 | 5.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 135.23 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 137.11 is outside the allowed range [80.00, 120.00] | FALSE | 006CALS.d | 2023-11-16 11:40 | CalStd | 3 | 1 ppb m16 + Al | 2.847 | 3.6 | 1.047 | 3 | <0.000 | N/A | <0.000 | N/A | 0.999 | 3.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 158.50 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 166.42 is outside the allowed range [80.00, 120.00] | FALSE | 007CALS.d | 2023-11-16 11:42 | CalStd | 4 | 5 ppb m16 + Al | 2.904 | 4 | 5.433 | 2.6 | 4.015 | 2.7 | 4.327 | 3.8 | 5.038 | 0.9
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 191.39 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 202.01 is outside the allowed range [80.00, 120.00] | FALSE | 008CALS.d | 2023-11-16 11:43 | CalStd | 5 | 10 ppb m16 + Al | 49.82 | 2.6 | 10.731 | 2.5 | 9.689 | 2.2 | 9.603 | 9.8 | 10.054 | 2
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 226.75 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 246.55 is outside the allowed range [80.00, 120.00] | FALSE | 009CALS.d | 2023-11-16 11:45 | CalStd | 6 | 50 ppb m16 + Al | 50.403 | 2.3 | 54.312 | 1.9 | 52.543 | 1.3 | 51.522 | 1.1 | 50.413 | 1.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 294.14 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 306.45 is outside the allowed range [80.00, 120.00] | FALSE | 010CALS.d | 2023-11-16 11:47 | CalStd | 7 | 100 ppb m16 + Al | 101.095 | 1.7 | 108.717 | 1.3 | 104.811 | 1.4 | 102.531 | 4.2 | 100.807 | 1.4
45 Sc ( ISTD ) [ He ] : CPS RSD value = 6.56 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 386.55 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 404.89 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 138.87 is outside the allowed range [80.00, 120.00] | FALSE | 011CALS.d | 2023-11-16 11:48 | CalStd | 8 | 500 ppb m16 + Al | 507.689 | 1.2 | 514.227 | 1.1 | 506.503 | 1 | 507.291 | 1.1 | 499.973 | 0.9
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 461.85 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 506.89 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 155.90 is outside the allowed range [80.00, 120.00] | FALSE | 012CALS.d | 2023-11-16 11:50 | CalStd | 9 | 1000 ppb m16 + Al | 996.166 | 1.9 | 991.79 | 2.2 | 996.15 | 2.3 | 996.034 | 2.4 | 999.911 | 1.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7414.54 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7507.11 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1229.69 is outside the allowed range [80.00, 120.00] | FALSE | 080SMPL.d | 2023-11-16 14:17 | CalStd | 10 | level 10 m6 | 1125.195 | 1.5 | 1121.019 | 1.9 | 10282.827 | 1.5 | 10045.684 | 1.1 | 10611.155 | 1.1
| TRUE | 081SMPL.d | 2023-11-16 14:19 | CalStd | 11 | level 11 m6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 562.55 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 5.07 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 584.98 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 159.39 is outside the allowed range [80.00, 120.00] | FALSE | 013SMPL.d | 2023-11-16 11:52 | Sample |  | blk | <0.000 | N/A | <0.000 | N/A | <0.000 | N/A | <0.000 | N/A | 0.052 | 3.2
45 Sc ( ISTD ) [ He ] : CPS RSD value = 7.74 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 628.50 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 5.34 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 671.37 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 175.99 is outside the allowed range [80.00, 120.00] | FALSE | 014SMPL.d | 2023-11-16 11:53 | Sample |  | m6 x100 x100 | <0.000 | N/A | 1.099 | 5.5 | 9.547 | 1.7 | 9.983 | 4.3 | 10 | 1.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 1644.32 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 1708.29 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 333.21 is outside the allowed range [80.00, 120.00] | FALSE | 015SMPL.d | 2023-11-16 12:13 | Sample |  | blk | <0.000 | N/A | 0.172 | 300.1 | <0.000 | N/A | <0.000 | N/A | <0.000 | N/A
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 1810.46 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 1886.67 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 462.93 is outside the allowed range [80.00, 120.00] | FALSE | 016SMPL.d | 2023-11-16 12:15 | Sample |  | 0-10-T | 277.27 | 0.3 | 78.506 | 2.4 | 745.162 | 0.4 | 746.449 | 1.4 | 1.439 | 1.4
56 Fe [ He ] : Concentration value = 1741.61 is over the calibration range57 Fe [ He ] : Concentration value = 1741.26 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 1914.26 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 1993.18 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 494.39 is outside the allowed range [80.00, 120.00] | FALSE | 017SMPL.d | 2023-11-16 12:16 | Sample |  | 0-20-T | 424.773 | 1.5 | 212.167 | 1.2 | 1741.609 | 2.2 | 1741.262 | 1.7 | 3.25 | 2.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2032.78 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2089.73 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 483.90 is outside the allowed range [80.00, 120.00] | FALSE | 018SMPL.d | 2023-11-16 12:18 | Sample |  | 0-30-T | 180.556 | 1.1 | 14.523 | 2.7 | 337.298 | 1.4 | 335.967 | 1.5 | 0.376 | 4.2
55 Mn [ He ] : CPS RSD value = 5.14 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2053.05 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2146.26 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 503.56 is outside the allowed range [80.00, 120.00] | FALSE | 019SMPL.d | 2023-11-16 12:20 | Sample |  | 0-40-T | 346.946 | 0.5 | 9.287 | 5.2 | 414.774 | 0.7 | 413.384 | 1.2 | 0.593 | 1.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2239.60 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2309.24 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 572.14 is outside the allowed range [80.00, 120.00] | FALSE | 020SMPL.d | 2023-11-16 12:21 | Sample |  | 0-60-T | 236.222 | 1.2 | 3.801 | 3.1 | 208.707 | 2.7 | 207.527 | 2.9 | 0.745 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2330.86 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2409.89 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 564.71 is outside the allowed range [80.00, 120.00] | FALSE | 021SMPL.d | 2023-11-16 12:23 | Sample |  | 0-70-T | 187.225 | 1.3 | 4.421 | 5.1 | 743.053 | 0.8 | 740.213 | 0.3 | 1.242 | 4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2425.36 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2513.06 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 572.13 is outside the allowed range [80.00, 120.00] | FALSE | 022SMPL.d | 2023-11-16 12:24 | Sample |  | 0-80-T | 139.274 | 0.7 | 4.448 | 2.3 | 817.19 | 0.5 | 815.517 | 1.1 | 1.023 | 2.9
56 Fe [ He ] : Concentration value = 1291.46 is over the calibration range57 Fe [ He ] : Concentration value = 1303.08 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2571.95 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2651.99 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 638.54 is outside the allowed range [80.00, 120.00] | FALSE | 023SMPL.d | 2023-11-16 12:26 | Sample |  | 0-90-T | 541.516 | 1.8 | 74.482 | 2.4 | 1291.463 | 1.3 | 1303.075 | 2.2 | 2.125 | 2.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2693.95 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2788.15 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 623.24 is outside the allowed range [80.00, 120.00] | FALSE | 024SMPL.d | 2023-11-16 12:28 | Sample |  | 0-100-T | 320.592 | 0.9 | 176.827 | 2 | 788.642 | 1.2 | 790.752 | 1 | 1.243 | 1.6
27 Al [ He ] : Concentration value = 1732.27 is over the calibration range56 Fe [ He ] : Concentration value = 1095.97 is over the calibration range57 Fe [ He ] : Concentration value = 1105.22 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2874.67 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2937.20 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 647.29 is outside the allowed range [80.00, 120.00] | FALSE | 025SMPL.d | 2023-11-16 12:29 | Sample |  | 0-10-B | 1732.265 | 0.7 | 148.667 | 1.9 | 1095.967 | 1.1 | 1105.216 | 1.1 | 3.766 | 2.1
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2882.46 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2970.87 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 597.55 is outside the allowed range [80.00, 120.00] | FALSE | 026SMPL.d | 2023-11-16 12:31 | Sample |  | blk | 2.288 | 5.9 | <0.000 | N/A | 2.55 | 5.7 | 2.809 | 2.9 | 0.002 | 75.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3023.36 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3111.09 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 591.79 is outside the allowed range [80.00, 120.00] | FALSE | 027SMPL.d | 2023-11-16 12:33 | Sample |  | 100 ppb check | 103.752 | 1.3 | 108.39 | 1.3 | 109.375 | 2.1 | 106.833 | 1.9 | 101.892 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3163.18 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3278.17 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 805.30 is outside the allowed range [80.00, 120.00] | FALSE | 028SMPL.d | 2023-11-16 12:34 | Sample |  | 0-50-T | 111.496 | 0.2 | 8.011 | 3.5 | 283.137 | 0.5 | 276.931 | 1.4 | 0.359 | 2.5
55 Mn [ He ] : CPS RSD value = 6.21 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3302.40 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3396.95 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 701.44 is outside the allowed range [80.00, 120.00] | FALSE | 029SMPL.d | 2023-11-16 12:36 | Sample |  | 0-60-B | 209.874 | 0.3 | 2.65 | 6.6 | 109.83 | 0.4 | 107.82 | 1.3 | 0.609 | 1.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3444.40 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3578.61 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 752.54 is outside the allowed range [80.00, 120.00] | FALSE | 030SMPL.d | 2023-11-16 12:37 | Sample |  | 0-100-B | 268.391 | 1 | 87.707 | 3.3 | 541.108 | 0.5 | 546.32 | 0.5 | 0.661 | 1.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3420.75 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3638.38 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 721.53 is outside the allowed range [80.00, 120.00] | FALSE | 031SMPL.d | 2023-11-16 12:39 | Sample |  | 0-70-B | 65.34 | 1.7 | 2.951 | 3.4 | 31.389 | 4.3 | 31.565 | 3.7 | 0.251 | 4.3
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3707.78 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3799.75 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 782.70 is outside the allowed range [80.00, 120.00] | FALSE | 032SMPL.d | 2023-11-16 12:41 | Sample |  | 20-100-T | 561.169 | 1.5 | 122.422 | 1.9 | 612.197 | 1.4 | 612.744 | 2.3 | 2.22 | 0.7
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3853.24 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3962.39 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 870.51 is outside the allowed range [80.00, 120.00] | FALSE | 033SMPL.d | 2023-11-16 12:42 | Sample |  | 20-80-T | 489.794 | 0.3 | 26.536 | 3.2 | 648.152 | 1.1 | 653.872 | 0.9 | 2.095 | 0.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4028.42 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 4159.61 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 908.52 is outside the allowed range [80.00, 120.00] | FALSE | 034SMPL.d | 2023-11-16 12:44 | Sample |  | 20-30-B | 512.147 | 0.8 | 103.814 | 3 | 882.246 | 0.5 | 882.337 | 0.8 | 4.931 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4145.75 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 4277.49 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 877.05 is outside the allowed range [80.00, 120.00] | FALSE | 035SMPL.d | 2023-11-16 12:46 | Sample |  | 20-70-T | 472.91 | 0.3 | 109.266 | 1.8 | 837.77 | 0.8 | 837.914 | 0.5 | 1.472 | 1.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4216.97 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 4291.13 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 859.16 is outside the allowed range [80.00, 120.00] | FALSE | 036SMPL.d | 2023-11-16 12:47 | Sample |  | 20-60-T | 336.476 | 0.3 | 78.308 | 2.3 | 785.85 | 0.2 | 782.193 | 0.6 | 3.523 | 0.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4913.47 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5096.65 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1057.18 is outside the allowed range [80.00, 120.00] | FALSE | 037SMPL.d | 2023-11-16 13:01 | Sample |  | 20-90-B | 569.005 | 1.3 | 39.096 | 2.8 | 365.807 | 1.1 | 370.425 | 1.9 | 1.187 | 2.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4908.89 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5123.18 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1284.40 is outside the allowed range [80.00, 120.00] | FALSE | 038SMPL.d | 2023-11-16 13:03 | Sample |  | blk | 1.119 | 24.3 | <0.000 | N/A | 2.444 | 5.8 | 2.421 | 15 | <0.000 | N/A
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5027.93 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5240.86 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1016.00 is outside the allowed range [80.00, 120.00] | FALSE | 039SMPL.d | 2023-11-16 13:05 | Sample |  | 100 ppb check | 105.769 | 2.7 | 110.052 | 1.8 | 111.561 | 2.6 | 106.407 | 3.6 | 104.647 | 2.1
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5238.40 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5433.09 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1109.07 is outside the allowed range [80.00, 120.00] | FALSE | 040SMPL.d | 2023-11-16 13:06 | Sample |  | 0-50-B | 307.673 | 0.5 | 1.108 | 5.1 | 228.903 | 1.2 | 228.437 | 0.4 | 0.361 | 2.9
27 Al [ He ] : Concentration value = 2359.84 is over the calibration range56 Fe [ He ] : Concentration value = 11436.07 is over the calibration range57 Fe [ He ] : Concentration value = 10905.77 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5325.07 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5495.62 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1135.33 is outside the allowed range [80.00, 120.00] | FALSE | 041SMPL.d | 2023-11-16 13:08 | Sample |  | 0-90-B | 2359.843 | 3 | 143.788 | 1.6 | 11436.075 | 2.5 | 10905.767 | 3 | 30.81 | 3.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5310.04 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5550.75 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1152.34 is outside the allowed range [80.00, 120.00] | FALSE | 042SMPL.d | 2023-11-16 13:10 | Sample |  | 20-40-T | 360.617 | 0.4 | 19.874 | 2.3 | 951.492 | 1.3 | 950.955 | 0.7 | 3.32 | 2
27 Al [ He ] : CPS RSD value = 17.60 is over the allowed maximum = 5.0056 Fe [ He ] : Concentration value = 1674.38 is over the calibration range56 Fe [ He ] : CPS RSD value = 19.13 is over the allowed maximum = 5.0057 Fe [ He ] : Concentration value = 1667.63 is over the calibration range57 Fe [ He ] : CPS RSD value = 19.60 is over the allowed maximum = 5.0075 As [ He ] : CPS RSD value = 19.56 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : CPS RSD value = 47.87 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5971.41 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 48.99 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6145.86 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 44.93 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1325.91 is outside the allowed range [80.00, 120.00] | FALSE | 043SMPL.d | 2023-11-16 13:11 | Sample |  | 20-20-T | 649.402 | 17.7 | 221.5 | 3.9 | 1674.385 | 19.2 | 1667.626 | 19.6 | 4.555 | 19.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5434.87 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 4681.69 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 877.07 is outside the allowed range [80.00, 120.00] | FALSE | 044SMPL.d | 2023-11-16 13:13 | Sample |  | 20-100-B | 869.815 | 0.8 | 69.78 | 3.3 | 460.542 | 0.5 | 457.66 | 0.7 | 3.81 | 1.6
27 Al [ He ] : Concentration value = 3105.11 is over the calibration range56 Fe [ He ] : Concentration value = 8862.93 is over the calibration range57 Fe [ He ] : Concentration value = 8620.22 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6191.64 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5800.87 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1159.77 is outside the allowed range [80.00, 120.00] | FALSE | 045SMPL.d | 2023-11-16 13:15 | Sample |  | 20-10-B | 3105.114 | 1 | 42.115 | 1.4 | 8862.928 | 0.6 | 8620.22 | 0.9 | 3.874 | 0.8
56 Fe [ He ] : Concentration value = 2817.90 is over the calibration range57 Fe [ He ] : Concentration value = 2801.82 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6235.70 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5924.56 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1115.63 is outside the allowed range [80.00, 120.00] | FALSE | 046SMPL.d | 2023-11-16 13:16 | Sample |  | 20-30-T | 649.932 | 1.2 | 292.989 | 1.7 | 2817.902 | 1 | 2801.82 | 0.8 | 5.559 | 0.4
27 Al [ He ] : Concentration value = 1029.05 is over the calibration range56 Fe [ He ] : Concentration value = 4017.39 is over the calibration range57 Fe [ He ] : Concentration value = 3989.62 is over the calibration range45 Sc ( ISTD ) [ He ] : CPS RSD value = 11.33 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5698.03 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 10.79 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5561.24 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 79.79 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1962.34 is outside the allowed range [80.00, 120.00] | FALSE | 047SMPL.d | 2023-11-16 13:18 | Sample |  | 10-0-T | 1029.048 | 1.6 | 759.059 | 3.6 | 4017.391 | 1.9 | 3989.623 | 1.3 | 6.067 | 1.7
27 Al [ He ] : Concentration value = 1519.20 is over the calibration range55 Mn [ He ] : Concentration value = 1365.58 is over the calibration range56 Fe [ He ] : Concentration value = 6070.51 is over the calibration range57 Fe [ He ] : Concentration value = 5949.81 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6306.07 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6026.19 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1169.80 is outside the allowed range [80.00, 120.00] | FALSE | 048SMPL.d | 2023-11-16 13:20 | Sample |  | 10-10-T | 1519.202 | 0.5 | 1365.579 | 1 | 6070.515 | 0.3 | 5949.807 | 0.7 | 14.52 | 1.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6314.62 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6081.04 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1147.08 is outside the allowed range [80.00, 120.00] | FALSE | 049SMPL.d | 2023-11-16 13:21 | Sample |  | 10-20-T | 139.711 | 0.9 | 202.726 | 3 | 314.475 | 0.6 | 312.171 | 0.6 | 0.671 | 1
56 Fe [ He ] : CPS RSD value = 6.07 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6288.10 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6060.34 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1104.28 is outside the allowed range [80.00, 120.00] | FALSE | 050SMPL.d | 2023-11-16 13:28 | Sample |  | blk | <0.000 | N/A | 0.141 | 12.5 | 2.723 | 11.8 | 3.668 | 15.2 | 0.001 | 270.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6344.41 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6135.81 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1105.24 is outside the allowed range [80.00, 120.00] | FALSE | 051SMPL.d | 2023-11-16 13:30 | Sample |  | 100 ppb check | 119.868 | 1.7 | 117.385 | 1.3 | 115.248 | 1.5 | 111.307 | 2.6 | 108.444 | 1.3
56 Fe [ He ] : Concentration value = 1359.59 is over the calibration range57 Fe [ He ] : Concentration value = 1353.11 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6531.51 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 6.13 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6449.18 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1244.12 is outside the allowed range [80.00, 120.00] | FALSE | 052SMPL.d | 2023-11-16 13:31 | Sample |  | 10-30-T | 501.987 | 0.7 | 151.618 | 3.2 | 1359.589 | 1.2 | 1353.109 | 1.1 | 3.356 | 1
27 Al [ He ] : Concentration value = 1555.17 is over the calibration range56 Fe [ He ] : Concentration value = 5942.02 is over the calibration range57 Fe [ He ] : Concentration value = 5855.82 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6684.39 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6881.45 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1212.20 is outside the allowed range [80.00, 120.00] | FALSE | 053SMPL.d | 2023-11-16 13:33 | Sample |  | 10-40-T | 1555.168 | 1 | 590.192 | 1.7 | 5942.02 | 1.1 | 5855.823 | 0.2 | 17.54 | 0.8
56 Fe [ He ] : Concentration value = 6262.98 is over the calibration range57 Fe [ He ] : Concentration value = 6196.11 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6709.82 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 7.50 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6597.05 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1267.72 is outside the allowed range [80.00, 120.00] | FALSE | 054SMPL.d | 2023-11-16 13:35 | Sample |  | 10-50-T | 735.946 | 1.3 | 998.657 | 1.5 | 6262.981 | 0.9 | 6196.11 | 0.2 | 15.661 | 0.2
56 Fe [ He ] : Concentration value = 1860.59 is over the calibration range57 Fe [ He ] : Concentration value = 1830.78 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6639.84 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 8.34 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6671.67 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1285.61 is outside the allowed range [80.00, 120.00] | FALSE | 055SMPL.d | 2023-11-16 13:36 | Sample |  | 10-80-T | 650.665 | 2.1 | 450.387 | 3.6 | 1860.588 | 1.8 | 1830.781 | 1.3 | 4.768 | 1.3
56 Fe [ He ] : Concentration value = 1473.90 is over the calibration range57 Fe [ He ] : Concentration value = 1453.39 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6799.58 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7007.90 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1233.16 is outside the allowed range [80.00, 120.00] | FALSE | 056SMPL.d | 2023-11-16 13:38 | Sample |  | 10-100-T | 638.533 | 0.6 | 288.95 | 1 | 1473.904 | 0.5 | 1453.386 | 1 | 3.092 | 1.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7052.25 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6848.77 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1093.77 is outside the allowed range [80.00, 120.00] | FALSE | 057SMPL.d | 2023-11-16 13:39 | Sample |  | 10-90-T | 32.5 | 2 | 128.963 | 0.8 | 40.28 | 3 | 41.502 | 3.8 | 0.146 | 2.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7042.88 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7287.33 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1280.35 is outside the allowed range [80.00, 120.00] | FALSE | 058SMPL.d | 2023-11-16 13:41 | Sample |  | 10-60-T | 538.181 | 0.5 | 133.456 | 1.9 | 518.778 | 0.5 | 505.286 | 0.9 | 5.645 | 1.4
56 Fe [ He ] : Concentration value = 6531.48 is over the calibration range57 Fe [ He ] : Concentration value = 6376.77 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7093.41 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7390.53 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1287.36 is outside the allowed range [80.00, 120.00] | FALSE | 059SMPL.d | 2023-11-16 13:43 | Sample |  | 10-70-T | 562.972 | 1.8 | 451.098 | 1.2 | 6531.483 | 1.4 | 6376.77 | 3 | 14.563 | 2.4
56 Fe [ He ] : Concentration value = 1332.52 is over the calibration range57 Fe [ He ] : Concentration value = 1321.40 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7113.01 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7333.62 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1363.90 is outside the allowed range [80.00, 120.00] | FALSE | 060SMPL.d | 2023-11-16 13:44 | Sample |  | 10-0-B | 370.028 | 0.3 | 583.919 | 1 | 1332.516 | 0.4 | 1321.395 | 0.4 | 2.714 | 0.9
56 Fe [ He ] : Concentration value = 1215.18 is over the calibration range57 Fe [ He ] : Concentration value = 1195.98 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7242.19 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7481.51 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1349.44 is outside the allowed range [80.00, 120.00] | FALSE | 061SMPL.d | 2023-11-16 13:46 | Sample |  | 10-20-B | 470.395 | 1.2 | 529.49 | 3.6 | 1215.184 | 1.4 | 1195.985 | 1.2 | 3.162 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6944.93 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7283.70 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1258.54 is outside the allowed range [80.00, 120.00] | FALSE | 062SMPL.d | 2023-11-16 13:48 | Sample |  | blk | 1.419 | 23.5 | 0.647 | 8.4 | 9.961 | 1.8 | 10.353 | 10.9 | 0.017 | 13.2
45 Sc ( ISTD ) [ He ] : CPS RSD value = 28.07 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7086.50 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 32.04 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7267.05 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1224.44 is outside the allowed range [80.00, 120.00] | FALSE | 063SMPL.d | 2023-11-16 13:49 | Sample |  | 100 ppb check | 122.344 | 4.8 | 118.312 | 4.7 | 122.415 | 4.3 | 117.712 | 4.4 | 108.443 | 4.5
27 Al [ He ] : Concentration value = 2513.93 is over the calibration range55 Mn [ He ] : Concentration value = 1342.39 is over the calibration range56 Fe [ He ] : Concentration value = 4676.06 is over the calibration range57 Fe [ He ] : Concentration value = 4626.66 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7352.98 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7572.16 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1374.75 is outside the allowed range [80.00, 120.00] | FALSE | 064SMPL.d | 2023-11-16 13:51 | Sample |  | 10-80-B | 2513.929 | 1.4 | 1342.39 | 2 | 4676.059 | 1.1 | 4626.664 | 1.3 | 14.335 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7267.86 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7535.41 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1256.32 is outside the allowed range [80.00, 120.00] | FALSE | 065SMPL.d | 2023-11-16 13:53 | Sample |  | 10-90-B | 95.038 | 2.9 | 123.769 | 1.6 | 106.473 | 1 | 103.618 | 1.7 | 0.309 | 1.6
27 Al [ He ] : Concentration value = 1030.98 is over the calibration range55 Mn [ He ] : Concentration value = 2156.99 is over the calibration range56 Fe [ He ] : Concentration value = 4101.93 is over the calibration range57 Fe [ He ] : Concentration value = 4053.28 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7295.68 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7665.68 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1370.39 is outside the allowed range [80.00, 120.00] | FALSE | 066SMPL.d | 2023-11-16 13:54 | Sample |  | 10-10-B | 1030.979 | 1.4 | 2156.988 | 1.6 | 4101.925 | 1.1 | 4053.278 | 1.5 | 11.978 | 0.7
27 Al [ He ] : Concentration value = 1053.91 is over the calibration range56 Fe [ He ] : Concentration value = 2660.51 is over the calibration range57 Fe [ He ] : Concentration value = 2598.34 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7425.78 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7705.62 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1362.52 is outside the allowed range [80.00, 120.00] | FALSE | 067SMPL.d | 2023-11-16 13:56 | Sample |  | 10-50-B | 1053.915 | 1.3 | 316.53 | 3.4 | 2660.508 | 1 | 2598.337 | 1.5 | 22.053 | 1.1
27 Al [ He ] : Concentration value = 2629.94 is over the calibration range56 Fe [ He ] : Concentration value = 6956.29 is over the calibration range57 Fe [ He ] : Concentration value = 6792.78 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7435.64 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7723.01 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1369.50 is outside the allowed range [80.00, 120.00] | FALSE | 068SMPL.d | 2023-11-16 13:58 | Sample |  | 10-60-B | 2629.944 | 1 | 207.922 | 1.8 | 6956.293 | 1.2 | 6792.781 | 0.8 | 45.529 | 0.9
56 Fe [ He ] : Concentration value = 2099.62 is over the calibration range57 Fe [ He ] : Concentration value = 2042.59 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7459.60 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7760.73 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1387.89 is outside the allowed range [80.00, 120.00] | FALSE | 069SMPL.d | 2023-11-16 13:59 | Sample |  | 10-70-B | 595.11 | 1.3 | 266.033 | 2.6 | 2099.615 | 1.2 | 2042.59 | 1.2 | 4.809 | 2
56 Fe [ He ] : Concentration value = 1163.82 is over the calibration range57 Fe [ He ] : Concentration value = 1138.44 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7398.22 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7704.86 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1331.05 is outside the allowed range [80.00, 120.00] | FALSE | 070SMPL.d | 2023-11-16 14:01 | Sample |  | 10-100-B | 519.842 | 1.3 | 138.095 | 0.6 | 1163.817 | 0.9 | 1138.44 | 0.9 | 2.62 | 2.7
27 Al [ He ] : Concentration value = 1061.86 is over the calibration range56 Fe [ He ] : Concentration value = 3463.68 is over the calibration range57 Fe [ He ] : Concentration value = 3382.74 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7530.53 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7813.42 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 7.89 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1342.41 is outside the allowed range [80.00, 120.00] | FALSE | 071SMPL.d | 2023-11-16 14:02 | Sample |  | 20-10-T | 1061.862 | 1.1 | 33.758 | 4.7 | 3463.684 | 1.1 | 3382.737 | 0.8 | 1.823 | 2.3
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7490.31 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7801.53 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 8.09 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1338.04 is outside the allowed range [80.00, 120.00] | FALSE | 072SMPL.d | 2023-11-16 14:04 | Sample |  | 20-50-T | 593.959 | 1.7 | 50.5 | 4.1 | 777.196 | 1.6 | 757.693 | 0.8 | 7.056 | 2.7
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7462.05 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7627.01 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1379.16 is outside the allowed range [80.00, 120.00] | FALSE | 073SMPL.d | 2023-11-16 14:06 | Sample |  | 20-70-B | 350.992 | 1.1 | 14.336 | 1.7 | 313.271 | 0.9 | 303.766 | 0.1 | 0.949 | 1.1
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7301.11 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7719.96 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1245.40 is outside the allowed range [80.00, 120.00] | FALSE | 074SMPL.d | 2023-11-16 14:07 | Sample |  | 20-60-B | 243.353 | 3.1 | 17.246 | 3.9 | 241.877 | 2.2 | 235.688 | 2.8 | 4.9 | 0.9
55 Mn [ He ] : CPS RSD value = 5.87 is over the allowed maximum = 5.0057 Fe [ He ] : CPS RSD value = 5.90 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7352.19 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7743.09 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1293.06 is outside the allowed range [80.00, 120.00] | FALSE | 075SMPL.d | 2023-11-16 14:09 | Sample |  | 20-80-B | 210.836 | 1.7 | 7.852 | 6 | 31.126 | 2.4 | 28.79 | 6.5 | 0.335 | 1.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7395.06 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7712.15 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1324.95 is outside the allowed range [80.00, 120.00] | FALSE | 076SMPL.d | 2023-11-16 14:11 | Sample |  | 20-90-T | 382.683 | 1.3 | 44.515 | 3 | 394.721 | 0.4 | 383.964 | 2.6 | 0.929 | 1.9
55 Mn [ He ] : CPS RSD value = 5.63 is over the allowed maximum = 5.0056 Fe [ He ] : CPS RSD value = 5.79 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7357.47 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7868.70 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1282.07 is outside the allowed range [80.00, 120.00] | FALSE | 077SMPL.d | 2023-11-16 14:12 | Sample |  | DI H2O | <0.000 | N/A | 1.321 | 6.3 | 4.593 | 9.1 | 6.053 | 15.2 | 0.018 | 14.8
56 Fe [ He ] : CPS RSD value = 5.53 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7403.16 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7762.21 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1401.26 is outside the allowed range [80.00, 120.00] | FALSE | 078SMPL.d | 2023-11-16 14:14 | Sample |  | blk | 4.528 | 2.9 | 0.629 | 4.5 | 10.329 | 6.9 | 10.679 | 2.7 | 0.221 | 5.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7481.82 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7716.10 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1231.85 is outside the allowed range [80.00, 120.00] | FALSE | 079SMPL.d | 2023-11-16 14:16 | Sample |  | 100 ppb check | 121.76 | 2.8 | 115.993 | 2.2 | 119.984 | 3.8 | 116.174 | 4.5 | 106.301 | 2.2
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7345.47 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 8102.91 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 9.17 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1373.47 is outside the allowed range [80.00, 120.00] | FALSE | 082SMPL.d | 2023-11-16 14:21 | Sample |  | blk | <0.000 | N/A | 0.712 | 4.3 | 11.066 | 2.8 | 11.555 | 6.8 | 10.087 | 2.2
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7331.14 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 8208.94 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1293.48 is outside the allowed range [80.00, 120.00] | FALSE | 083SMPL.d | 2023-11-16 14:22 | Sample |  | blk | <0.000 | N/A | 0.325 | 2.3 | 5.412 | 5.6 | 6.044 | 8.7 | 3.781 | 1.2
56 Fe [ He ] : CPS RSD value = 5.26 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7408.82 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 8056.24 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1241.47 is outside the allowed range [80.00, 120.00] | FALSE | 084SMPL.d | 2023-11-16 14:24 | Sample |  | blk | <0.000 | N/A | 0.136 | 11.4 | 3.316 | 9.4 | 3.153 | 13.2 | 2.206 | 0.8