## ICP_Mn_As_Calib
Sample |  |  |  |  | 55 Mn [ He ] |  |  | 75 As [ He ] |  |  | 45 Sc ( ISTD ) [ He ] |  | 103 Rh ( ISTD ) [ He ] |  | 193 Ir ( ISTD ) [ He ] |  | 55 Mn [ He ] |  | 75 As [ He ]
| Type | Level | Sample Name | conc_calib_ppb | CPS | CPS RSD | Mn_conc_sample_ppb | CPS | CPS RSD | As_conc_sample_ppb | CPS | CPS RSD | CPS | CPS RSD | CPS | CPS RSD | Conc. [ ppb ] | Conc. RSD | Conc. [ ppb ] | Conc. RSD
| Sample |  | blk |  | 1360.11 | 3.2 |  | 13 | 23.1 |  | 33.33 | 69.3 | 110.01 | 0 | 590.03 | 12.8
56 Fe [ He ] : CPS RSD value = 5.02 is over the allowed maximum = 5.00 | Sample |  | blk |  | 690.04 | 20.4 |  | 15.67 | 35.2 |  | 46.67 | 53.9 | 140 | 21.4 | 600.04 | 10.1
| Sample |  | blk |  | 383.36 | 15.9 |  | 8.67 | 6.7 |  | 6288.22 | 3.6 | 54600.29 | 0.4 | 1123.42 | 3.1
| CalBlk | 1 | 0 ppb | 0 | 1416.78 | 12.4 |  | 32.33 | 26.3 |  | 2707.03 | 9 | 22608.86 | 3.4 | 763.38 | 9.5 | 0 | N/A | 0 | N/A
| CalStd | 2 | 0.5 ppb m16 + Al | 0.5 | 5757.99 | 4.5 |  | 1315.4 | 5.6 |  | 2790.38 | 1.9 | 26562.51 | 1.3 | 773.39 | 7.9 | 0.48 | 6 | 0.498 | 5.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 135.23 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 137.11 is outside the allowed range [80.00, 120.00] | CalStd | 3 | 1 ppb m16 + Al | 1 | 10894.51 | 2.6 |  | 2604.24 | 3.4 |  | 3660.59 | 9.5 | 30998.35 | 3.4 | 760.05 | 29.7 | 1.047 | 3 | 0.999 | 3.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 158.50 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 166.42 is outside the allowed range [80.00, 120.00] | CalStd | 4 | 5 ppb m16 + Al | 5 | 50599.4 | 2.6 |  | 13001.57 | 0.9 |  | 4290.77 | 0.8 | 37624.87 | 1.9 | 796.73 | 3.8 | 5.433 | 2.6 | 5.038 | 0.9
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 191.39 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 202.01 is outside the allowed range [80.00, 120.00] | CalStd | 5 | 10 ppb m16 + Al | 10 | 98558.31 | 2.5 |  | 25914.34 | 2 |  | 5181.09 | 4.7 | 45672.01 | 1.6 | 800.06 | 12.1 | 10.731 | 2.5 | 10.054 | 2
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 226.75 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 246.55 is outside the allowed range [80.00, 120.00] | CalStd | 6 | 50 ppb m16 + Al | 50 | 493088.85 | 1.9 |  | 129814.73 | 1.4 |  | 6138.16 | 4.1 | 55741.53 | 3.2 | 750.05 | 10.1 | 54.312 | 1.9 | 50.413 | 1.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 294.14 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 306.45 is outside the allowed range [80.00, 120.00] | CalStd | 7 | 100 ppb m16 + Al | 100 | 985597.82 | 1.3 |  | 259551.1 | 1.4 |  | 7962.46 | 4.3 | 69284.83 | 3.1 | 886.73 | 16.2 | 108.717 | 1.3 | 100.807 | 1.4
45 Sc ( ISTD ) [ He ] : CPS RSD value = 6.56 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 386.55 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 404.89 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 138.87 is outside the allowed range [80.00, 120.00] | CalStd | 8 | 500 ppb m16 + Al | 500 | 4656554.62 | 1.1 |  | 1287165.08 | 0.9 |  | 10464.12 | 6.6 | 91541.18 | 3.1 | 1060.08 | 12.3 | 514.227 | 1.1 | 499.973 | 0.9
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 461.85 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 506.89 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 155.90 is outside the allowed range [80.00, 120.00] | CalStd | 9 | 1000 ppb m16 + Al | 1000 | 8979786.32 | 2.2 |  | 2574207.08 | 1.8 |  | 12502.48 | 3.9 | 114602.58 | 3.1 | 1190.1 | 8.1 | 991.79 | 2.2 | 999.911 | 1.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7414.54 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7507.11 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1229.69 is outside the allowed range [80.00, 120.00] | CalStd | 10 | level 10 m6 |  | 10149657.14 | 1.9 |  | 27317422.67 | 1.1 |  | 200714.18 | 2.3 | 1697271.9 | 1.7 | 9387.18 | 5.5 | 1121.019 | 1.9 | 10611.155 | 1.1
| CalStd | 11 | level 11 m6 |  | 88454685.35 | 0.8 |  | 249936965.33 | 0.7 |  | 180908.09 | 1.6 | 1348858.31 | 1.8 | 8643.16 | 2.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 562.55 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 5.07 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 584.98 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 159.39 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 850.06 | 18.9 |  | 167 | 2.6 |  | 15228.45 | 4.8 | 132257.44 | 5.1 | 1216.77 | 2.5 | <0.000 | N/A | 0.052 | 3.2
45 Sc ( ISTD ) [ He ] : CPS RSD value = 7.74 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 628.50 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 5.34 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 671.37 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 175.99 is outside the allowed range [80.00, 120.00] | Sample |  | m6 x100 x100 |  | 11368.11 | 4.8 |  | 25775.4 | 1.6 |  | 17013.83 | 7.7 | 151788.8 | 5.3 | 1343.46 | 2.3 | 1.099 | 5.5 | 10 | 1.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 1644.32 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 1708.29 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 333.21 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 2974.82 | 157.2 |  | 16.67 | 30.2 |  | 44512.22 | 2.8 | 386224.16 | 2.7 | 2543.67 | 7 | 0.172 | 300.1 | <0.000 | N/A
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 1810.46 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 1886.67 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 462.93 is outside the allowed range [80.00, 120.00] | Sample |  | 0-10-T |  | 712110.22 | 2.4 |  | 3737.17 | 1.4 |  | 49009.82 | 3.1 | 426554.58 | 2.9 | 3533.95 | 3.3 | 78.506 | 2.4 | 1.439 | 1.4
56 Fe [ He ] : Concentration value = 1741.61 is over the calibration range57 Fe [ He ] : Concentration value = 1741.26 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 1914.26 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 1993.18 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 494.39 is outside the allowed range [80.00, 120.00] | Sample |  | 0-20-T |  | 1922096.74 | 1.2 |  | 8398.14 | 2.4 |  | 51819.63 | 2.8 | 450635.99 | 2.8 | 3774.06 | 7.2 | 212.167 | 1.2 | 3.25 | 2.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2032.78 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2089.73 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 483.90 is outside the allowed range [80.00, 120.00] | Sample |  | 0-30-T |  | 132889.63 | 2.7 |  | 1000.37 | 4.1 |  | 55028.08 | 2.8 | 472463.03 | 4 | 3693.99 | 5.2 | 14.523 | 2.7 | 0.376 | 4.2
55 Mn [ He ] : CPS RSD value = 5.14 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2053.05 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2146.26 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 503.56 is outside the allowed range [80.00, 120.00] | Sample |  | 0-40-T |  | 85485.96 | 5.1 |  | 1559.76 | 1.4 |  | 55576.77 | 2.7 | 485245.67 | 2.7 | 3844.04 | 7 | 9.287 | 5.2 | 0.593 | 1.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2239.60 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2309.24 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 572.14 is outside the allowed range [80.00, 120.00] | Sample |  | 0-60-T |  | 35828.4 | 3 |  | 1950.47 | 1.5 |  | 60626.77 | 2.3 | 522091.83 | 1.3 | 4367.57 | 5.7 | 3.801 | 3.1 | 0.745 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2330.86 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2409.89 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 564.71 is outside the allowed range [80.00, 120.00] | Sample |  | 0-70-T |  | 41437.02 | 4.9 |  | 3229.7 | 4 |  | 63097.18 | 3.6 | 544848.33 | 1.4 | 4310.87 | 3 | 4.421 | 5.1 | 1.242 | 4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2425.36 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2513.06 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 572.13 is outside the allowed range [80.00, 120.00] | Sample |  | 0-80-T |  | 41684.29 | 2.2 |  | 2666.93 | 2.9 |  | 65655.21 | 3.6 | 568173.23 | 2.8 | 4367.54 | 5.7 | 4.448 | 2.3 | 1.023 | 2.9
56 Fe [ He ] : Concentration value = 1291.46 is over the calibration range57 Fe [ He ] : Concentration value = 1303.08 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2571.95 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2651.99 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 638.54 is outside the allowed range [80.00, 120.00] | Sample |  | 0-90-T |  | 675683.27 | 2.4 |  | 5504.07 | 2.4 |  | 69623.44 | 0.9 | 599584.25 | 2.1 | 4874.45 | 4.9 | 74.482 | 2.4 | 2.125 | 2.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2693.95 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2788.15 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 623.24 is outside the allowed range [80.00, 120.00] | Sample |  | 0-100-T |  | 1602179.87 | 2 |  | 3232.04 | 1.6 |  | 72926.19 | 1.5 | 630367.96 | 1.2 | 4757.67 | 10.1 | 176.827 | 2 | 1.243 | 1.6
27 Al [ He ] : Concentration value = 1732.27 is over the calibration range56 Fe [ He ] : Concentration value = 1095.97 is over the calibration range57 Fe [ He ] : Concentration value = 1105.22 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2874.67 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2937.20 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 647.29 is outside the allowed range [80.00, 120.00] | Sample |  | 0-10-B |  | 1347252.22 | 1.9 |  | 9727.64 | 2 |  | 77818.2 | 2 | 664066.34 | 0.9 | 4941.3 | 9.4 | 148.667 | 1.9 | 3.766 | 2.1
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 2882.46 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 2970.87 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 597.55 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 910.06 | 7.2 |  | 37.67 | 10.7 |  | 78029.21 | 3.8 | 671680.2 | 4.2 | 4561.59 | 7.4 | <0.000 | N/A | 0.002 | 75.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3023.36 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3111.09 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 591.79 is outside the allowed range [80.00, 120.00] | Sample |  | 100 ppb check |  | 982640.22 | 1.3 |  | 262344.45 | 1.5 |  | 81843.4 | 3.7 | 703381.97 | 1.3 | 4517.61 | 2.8 | 108.39 | 1.3 | 101.892 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3163.18 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3278.17 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 805.30 is outside the allowed range [80.00, 120.00] | Sample |  | 0-50-T |  | 73942.18 | 3.4 |  | 956.04 | 2.4 |  | 85628.41 | 1.9 | 741157.72 | 2.1 | 6147.5 | 16.2 | 8.011 | 3.5 | 0.359 | 2.5
55 Mn [ He ] : CPS RSD value = 6.21 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3302.40 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3396.95 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 701.44 is outside the allowed range [80.00, 120.00] | Sample |  | 0-60-B |  | 25406.39 | 6.2 |  | 1599.1 | 1.8 |  | 89397.19 | 3 | 768011.5 | 2.6 | 5354.66 | 1.1 | 2.65 | 6.6 | 0.609 | 1.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3444.40 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3578.61 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 752.54 is outside the allowed range [80.00, 120.00] | Sample |  | 0-100-B |  | 795398.92 | 3.3 |  | 1733.11 | 1.4 |  | 93241.1 | 2.6 | 809082.57 | 3.4 | 5744.76 | 6.2 | 87.707 | 3.3 | 0.661 | 1.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3420.75 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3638.38 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 721.53 is outside the allowed range [80.00, 120.00] | Sample |  | 0-70-B |  | 28131.54 | 3.2 |  | 679.35 | 4.1 |  | 92600.92 | 2.2 | 822596.97 | 2.2 | 5508.05 | 5.7 | 2.951 | 3.4 | 0.251 | 4.3
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3707.78 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3799.75 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 782.70 is outside the allowed range [80.00, 120.00] | Sample |  | 20-100-T |  | 1109670.17 | 1.9 |  | 5747.16 | 0.7 |  | 100370.84 | 2.6 | 859080.14 | 3.1 | 5974.95 | 4.2 | 122.422 | 1.9 | 2.22 | 0.7
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 3853.24 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 3962.39 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 870.51 is outside the allowed range [80.00, 120.00] | Sample |  | 20-80-T |  | 241642.72 | 3.1 |  | 5426.04 | 0.4 |  | 104308.38 | 4.2 | 895850.9 | 3 | 6645.27 | 5.4 | 26.536 | 3.2 | 2.095 | 0.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4028.42 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 4159.61 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 908.52 is outside the allowed range [80.00, 120.00] | Sample |  | 20-30-B |  | 941214.41 | 3 |  | 12727.99 | 1.4 |  | 109050.73 | 2.8 | 940440.27 | 3.5 | 6935.45 | 3.7 | 103.814 | 3 | 4.931 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4145.75 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 4277.49 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 877.05 is outside the allowed range [80.00, 120.00] | Sample |  | 20-70-T |  | 990565.4 | 1.8 |  | 3822.18 | 1.3 |  | 112226.79 | 2.5 | 967091.39 | 3.5 | 6695.25 | 3.8 | 109.266 | 1.8 | 1.472 | 1.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4216.97 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 4291.13 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 859.16 is outside the allowed range [80.00, 120.00] | Sample |  | 20-60-T |  | 710318.66 | 2.3 |  | 9102.23 | 0.6 |  | 114154.89 | 4.2 | 970176.49 | 3.4 | 6558.62 | 6.4 | 78.308 | 2.3 | 3.523 | 0.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4913.47 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5096.65 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1057.18 is outside the allowed range [80.00, 120.00] | Sample |  | 20-90-B |  | 355340.85 | 2.8 |  | 3088.67 | 2.3 |  | 133009.35 | 3.4 | 1152294.72 | 3.1 | 8070.27 | 11.9 | 39.096 | 2.8 | 1.187 | 2.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 4908.89 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5123.18 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1284.40 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 1013.4 | 10.6 |  | 26 | 3.8 |  | 132885.17 | 4.1 | 1158292.53 | 3.7 | 9804.85 | 39.8 | <0.000 | N/A | <0.000 | N/A
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5027.93 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5240.86 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1016.00 is outside the allowed range [80.00, 120.00] | Sample |  | 100 ppb check |  | 997681.26 | 1.8 |  | 269436.16 | 2.1 |  | 136107.75 | 3 | 1184898.99 | 2.6 | 7755.96 | 6.4 | 110.052 | 1.8 | 104.647 | 2.1
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5238.40 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5433.09 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1109.07 is outside the allowed range [80.00, 120.00] | Sample |  | 0-50-B |  | 11448.25 | 4.5 |  | 962.04 | 2.8 |  | 141805.27 | 2.9 | 1228359.36 | 2.3 | 8466.39 | 4.6 | 1.108 | 5.1 | 0.361 | 2.9
27 Al [ He ] : Concentration value = 2359.84 is over the calibration range56 Fe [ He ] : Concentration value = 11436.07 is over the calibration range57 Fe [ He ] : Concentration value = 10905.77 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5325.07 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5495.62 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1135.33 is outside the allowed range [80.00, 120.00] | Sample |  | 0-90-B |  | 1303082.09 | 1.6 |  | 79350.57 | 3.4 |  | 144151.39 | 1.9 | 1242497.59 | 2 | 8666.91 | 6.9 | 143.788 | 1.6 | 30.81 | 3.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5310.04 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5550.75 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1152.34 is outside the allowed range [80.00, 120.00] | Sample |  | 20-40-T |  | 181331.42 | 2.2 |  | 8579.91 | 2 |  | 143744.62 | 2.3 | 1254960.92 | 2.6 | 8796.77 | 3 | 19.874 | 2.3 | 3.32 | 2
27 Al [ He ] : CPS RSD value = 17.60 is over the allowed maximum = 5.0056 Fe [ He ] : Concentration value = 1674.38 is over the calibration range56 Fe [ He ] : CPS RSD value = 19.13 is over the allowed maximum = 5.0057 Fe [ He ] : Concentration value = 1667.63 is over the calibration range57 Fe [ He ] : CPS RSD value = 19.60 is over the allowed maximum = 5.0075 As [ He ] : CPS RSD value = 19.56 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : CPS RSD value = 47.87 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5971.41 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 48.99 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6145.86 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 44.93 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1325.91 is outside the allowed range [80.00, 120.00] | Sample |  | 20-20-T |  | 2006588.36 | 3.9 |  | 11757.64 | 19.6 |  | 161647.97 | 47.9 | 1389508.76 | 49 | 10121.76 | 44.9 | 221.5 | 3.9 | 4.555 | 19.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5434.87 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 4681.69 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 877.07 is outside the allowed range [80.00, 120.00] | Sample |  | 20-100-B |  | 633111.06 | 3.3 |  | 9841.39 | 1.6 |  | 147123.74 | 3.7 | 1058476.36 | 3.1 | 6695.35 | 3.9 | 69.78 | 3.3 | 3.81 | 1.6
27 Al [ He ] : Concentration value = 3105.11 is over the calibration range56 Fe [ He ] : Concentration value = 8862.93 is over the calibration range57 Fe [ He ] : Concentration value = 8620.22 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6191.64 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5800.87 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1159.77 is outside the allowed range [80.00, 120.00] | Sample |  | 20-10-B |  | 382672.59 | 1.4 |  | 10004.5 | 0.8 |  | 167609.71 | 2.2 | 1311511.23 | 2.1 | 8853.47 | 6.3 | 42.115 | 1.4 | 3.874 | 0.8
56 Fe [ He ] : Concentration value = 2817.90 is over the calibration range57 Fe [ He ] : Concentration value = 2801.82 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6235.70 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5924.56 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1115.63 is outside the allowed range [80.00, 120.00] | Sample |  | 20-30-T |  | 2653753.87 | 1.7 |  | 14344.51 | 0.4 |  | 168802.42 | 3.5 | 1339476.13 | 3 | 8516.53 | 4.4 | 292.989 | 1.7 | 5.559 | 0.4
27 Al [ He ] : Concentration value = 1029.05 is over the calibration range56 Fe [ He ] : Concentration value = 4017.39 is over the calibration range57 Fe [ He ] : Concentration value = 3989.62 is over the calibration range45 Sc ( ISTD ) [ He ] : CPS RSD value = 11.33 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 5698.03 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 10.79 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 5561.24 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 79.79 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1962.34 is outside the allowed range [80.00, 120.00] | Sample |  | 10-0-T |  | 6872950.11 | 3.6 |  | 15652.21 | 1.7 |  | 154247.58 | 11.3 | 1257332.06 | 10.8 | 14980.13 | 79.8 | 759.059 | 3.6 | 6.067 | 1.7
27 Al [ He ] : Concentration value = 1519.20 is over the calibration range55 Mn [ He ] : Concentration value = 1365.58 is over the calibration range56 Fe [ He ] : Concentration value = 6070.51 is over the calibration range57 Fe [ He ] : Concentration value = 5949.81 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6306.07 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6026.19 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1169.80 is outside the allowed range [80.00, 120.00] | Sample |  | 10-10-T |  | 12363584.82 | 1 |  | 37412.18 | 1.6 |  | 170707.37 | 1.5 | 1362452.43 | 2.6 | 8930.02 | 3.5 | 1365.579 | 1 | 14.52 | 1.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6314.62 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6081.04 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1147.08 is outside the allowed range [80.00, 120.00] | Sample |  | 10-20-T |  | 1836635.13 | 3 |  | 1759.11 | 0.9 |  | 170938.9 | 2.5 | 1374854.82 | 2.3 | 8756.6 | 5.6 | 202.726 | 3 | 0.671 | 1
56 Fe [ He ] : CPS RSD value = 6.07 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6288.10 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6060.34 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1104.28 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 2690.35 | 5.9 |  | 35 | 20.6 |  | 170221.05 | 2.6 | 1370174.62 | 3 | 8429.87 | 4.7 | 0.141 | 12.5 | 0.001 | 270.4
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6344.41 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6135.81 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1105.24 is outside the allowed range [80.00, 120.00] | Sample |  | 100 ppb check |  | 1064070.87 | 1.2 |  | 279209.91 | 1.3 |  | 171745.36 | 2.9 | 1387237.32 | 2 | 8437.18 | 3.3 | 117.385 | 1.3 | 108.444 | 1.3
56 Fe [ He ] : Concentration value = 1359.59 is over the calibration range57 Fe [ He ] : Concentration value = 1353.11 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6531.51 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 6.13 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6449.18 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1244.12 is outside the allowed range [80.00, 120.00] | Sample |  | 10-30-T |  | 1373970.76 | 3.2 |  | 8672.63 | 1 |  | 176810.27 | 3.6 | 1458086.75 | 6.1 | 9497.34 | 2.4 | 151.618 | 3.2 | 3.356 | 1
27 Al [ He ] : Concentration value = 1555.17 is over the calibration range56 Fe [ He ] : Concentration value = 5942.02 is over the calibration range57 Fe [ He ] : Concentration value = 5855.82 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6684.39 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6881.45 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1212.20 is outside the allowed range [80.00, 120.00] | Sample |  | 10-40-T |  | 5344242.52 | 1.7 |  | 45186.41 | 0.8 |  | 180948.75 | 3.2 | 1555817.11 | 2.6 | 9253.69 | 2 | 590.192 | 1.7 | 17.54 | 0.8
56 Fe [ He ] : Concentration value = 6262.98 is over the calibration range57 Fe [ He ] : Concentration value = 6196.11 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6709.82 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 7.50 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6597.05 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1267.72 is outside the allowed range [80.00, 120.00] | Sample |  | 10-50-T |  | 9041950.91 | 1.5 |  | 40350.76 | 0.2 |  | 181637.13 | 4 | 1491518.67 | 7.5 | 9677.55 | 5.4 | 998.657 | 1.5 | 15.661 | 0.2
56 Fe [ He ] : Concentration value = 1860.59 is over the calibration range57 Fe [ He ] : Concentration value = 1830.78 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6639.84 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 8.34 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6671.67 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1285.61 is outside the allowed range [80.00, 120.00] | Sample |  | 10-80-T |  | 4078635.35 | 3.6 |  | 12307.29 | 1.3 |  | 179742.8 | 4.4 | 1508387.89 | 8.3 | 9814.06 | 1.4 | 450.387 | 3.6 | 4.768 | 1.3
56 Fe [ He ] : Concentration value = 1473.90 is over the calibration range57 Fe [ He ] : Concentration value = 1453.39 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6799.58 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7007.90 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1233.16 is outside the allowed range [80.00, 120.00] | Sample |  | 10-100-T |  | 2617190.59 | 1 |  | 7992.58 | 1.6 |  | 184066.88 | 3.9 | 1584405.6 | 3.7 | 9413.72 | 4.7 | 288.95 | 1 | 3.092 | 1.6
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7052.25 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 6848.77 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1093.77 is outside the allowed range [80.00, 120.00] | Sample |  | 10-90-T |  | 1168880.42 | 0.8 |  | 409.01 | 2.6 |  | 190906.8 | 2.4 | 1548428.73 | 4.5 | 8349.65 | 3.3 | 128.963 | 0.8 | 0.146 | 2.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7042.88 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7287.33 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1280.35 is outside the allowed range [80.00, 120.00] | Sample |  | 10-60-T |  | 1209550.71 | 1.9 |  | 14564.4 | 1.4 |  | 190653.1 | 2.8 | 1647582.68 | 1.6 | 9773.94 | 6.4 | 133.456 | 1.9 | 5.645 | 1.4
56 Fe [ He ] : Concentration value = 6531.48 is over the calibration range57 Fe [ He ] : Concentration value = 6376.77 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7093.41 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7390.53 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1287.36 is outside the allowed range [80.00, 120.00] | Sample |  | 10-70-T |  | 4085069.31 | 1.2 |  | 37522.51 | 2.4 |  | 192020.93 | 1.3 | 1670914.14 | 0.6 | 9827.45 | 3.8 | 451.098 | 1.2 | 14.563 | 2.4
56 Fe [ He ] : Concentration value = 1332.52 is over the calibration range57 Fe [ He ] : Concentration value = 1321.40 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7113.01 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7333.62 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1363.90 is outside the allowed range [80.00, 120.00] | Sample |  | 10-0-B |  | 5287458.67 | 1 |  | 7019.4 | 0.9 |  | 192551.42 | 2.2 | 1658047.94 | 2.9 | 10411.76 | 4.8 | 583.919 | 1 | 2.714 | 0.9
56 Fe [ He ] : Concentration value = 1215.18 is over the calibration range57 Fe [ He ] : Concentration value = 1195.98 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7242.19 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7481.51 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1349.44 is outside the allowed range [80.00, 120.00] | Sample |  | 10-20-B |  | 4794726.8 | 3.6 |  | 8171.34 | 1.5 |  | 196048.52 | 2.1 | 1691485.24 | 2.6 | 10301.36 | 3.5 | 529.49 | 3.6 | 3.162 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 6944.93 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7283.70 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1258.54 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 7272.23 | 6.8 |  | 76.67 | 7.6 |  | 188001.7 | 3.8 | 1646761.64 | 4.2 | 9607.42 | 4.1 | 0.647 | 8.4 | 0.017 | 13.2
45 Sc ( ISTD ) [ He ] : CPS RSD value = 28.07 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7086.50 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : CPS RSD value = 32.04 is over the allowed maximum = 5.00103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7267.05 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1224.44 is outside the allowed range [80.00, 120.00] | Sample |  | 100 ppb check |  | 1072462.33 | 4.7 |  | 279207.51 | 4.5 |  | 191833.94 | 28.1 | 1642997.11 | 32 | 9347.14 | 21.7 | 118.312 | 4.7 | 108.443 | 4.5
27 Al [ He ] : Concentration value = 2513.93 is over the calibration range55 Mn [ He ] : Concentration value = 1342.39 is over the calibration range56 Fe [ He ] : Concentration value = 4676.06 is over the calibration range57 Fe [ He ] : Concentration value = 4626.66 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7352.98 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7572.16 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1374.75 is outside the allowed range [80.00, 120.00] | Sample |  | 10-80-B |  | 12153659.82 | 2 |  | 36935.3 | 1.5 |  | 199047.63 | 2.2 | 1711979.04 | 2.2 | 10494.6 | 2.2 | 1342.39 | 2 | 14.335 | 1.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7267.86 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7535.41 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1256.32 is outside the allowed range [80.00, 120.00] | Sample |  | 10-90-B |  | 1121863.68 | 1.6 |  | 828.36 | 1.5 |  | 196743.36 | 3.6 | 1703670.96 | 2.3 | 9590.47 | 2.1 | 123.769 | 1.6 | 0.309 | 1.6
27 Al [ He ] : Concentration value = 1030.98 is over the calibration range55 Mn [ He ] : Concentration value = 2156.99 is over the calibration range56 Fe [ He ] : Concentration value = 4101.93 is over the calibration range57 Fe [ He ] : Concentration value = 4053.28 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7295.68 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7665.68 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1370.39 is outside the allowed range [80.00, 120.00] | Sample |  | 10-10-B |  | 19527969.29 | 1.6 |  | 30868.39 | 0.7 |  | 197496.38 | 2.7 | 1733122.58 | 2.3 | 10461.31 | 3.8 | 2156.988 | 1.6 | 11.978 | 0.7
27 Al [ He ] : Concentration value = 1053.91 is over the calibration range56 Fe [ He ] : Concentration value = 2660.51 is over the calibration range57 Fe [ He ] : Concentration value = 2598.34 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7425.78 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7705.62 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1362.52 is outside the allowed range [80.00, 120.00] | Sample |  | 10-50-B |  | 2866866.31 | 3.4 |  | 56806.76 | 1.1 |  | 201018.26 | 2.1 | 1742151.74 | 2.3 | 10401.21 | 3.6 | 316.53 | 3.4 | 22.053 | 1.1
27 Al [ He ] : Concentration value = 2629.94 is over the calibration range56 Fe [ He ] : Concentration value = 6956.29 is over the calibration range57 Fe [ He ] : Concentration value = 6792.78 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7435.64 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7723.01 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1369.50 is outside the allowed range [80.00, 120.00] | Sample |  | 10-60-B |  | 1883667.26 | 1.8 |  | 117241.79 | 0.9 |  | 201285.28 | 3 | 1746084.45 | 2.7 | 10454.51 | 4.1 | 207.922 | 1.8 | 45.529 | 0.9
56 Fe [ He ] : Concentration value = 2099.62 is over the calibration range57 Fe [ He ] : Concentration value = 2042.59 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7459.60 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7760.73 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1387.89 is outside the allowed range [80.00, 120.00] | Sample |  | 10-70-B |  | 2409732.83 | 2.6 |  | 12412.71 | 2 |  | 201933.85 | 4.3 | 1754612.01 | 3.5 | 10594.9 | 2 | 266.033 | 2.6 | 4.809 | 2
56 Fe [ He ] : Concentration value = 1163.82 is over the calibration range57 Fe [ He ] : Concentration value = 1138.44 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7398.22 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7704.86 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1331.05 is outside the allowed range [80.00, 120.00] | Sample |  | 10-100-B |  | 1251545.58 | 0.6 |  | 6778.29 | 2.7 |  | 200272.37 | 1.6 | 1741981.95 | 2.9 | 10160.94 | 1.5 | 138.095 | 0.6 | 2.62 | 2.7
27 Al [ He ] : Concentration value = 1061.86 is over the calibration range56 Fe [ He ] : Concentration value = 3463.68 is over the calibration range57 Fe [ He ] : Concentration value = 3382.74 is over the calibration range45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7530.53 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7813.42 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 7.89 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1342.41 is outside the allowed range [80.00, 120.00] | Sample |  | 20-10-T |  | 307016.26 | 4.7 |  | 4724.46 | 2.3 |  | 203853.83 | 4.3 | 1766525.81 | 4.5 | 10247.67 | 7.9 | 33.758 | 4.7 | 1.823 | 2.3
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7490.31 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7801.53 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 8.09 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1338.04 is outside the allowed range [80.00, 120.00] | Sample |  | 20-50-T |  | 458574.08 | 4.1 |  | 18196.86 | 2.7 |  | 202765.11 | 4 | 1763836.85 | 4.1 | 10214.36 | 8.1 | 50.5 | 4.1 | 7.056 | 2.7
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7462.05 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7627.01 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1379.16 is outside the allowed range [80.00, 120.00] | Sample |  | 20-70-B |  | 131200.03 | 1.7 |  | 2474.89 | 1.1 |  | 202000.13 | 2.6 | 1724379.35 | 2.5 | 10528.22 | 1.8 | 14.336 | 1.7 | 0.949 | 1.1
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7301.11 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7719.96 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1245.40 is outside the allowed range [80.00, 120.00] | Sample |  | 20-60-B |  | 157542.82 | 3.9 |  | 12646.25 | 0.8 |  | 197643.47 | 2.9 | 1745395.86 | 1.2 | 9507.13 | 2.2 | 17.246 | 3.9 | 4.9 | 0.9
55 Mn [ He ] : CPS RSD value = 5.87 is over the allowed maximum = 5.0057 Fe [ He ] : CPS RSD value = 5.90 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7352.19 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7743.09 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1293.06 is outside the allowed range [80.00, 120.00] | Sample |  | 20-80-B |  | 72497.35 | 5.9 |  | 895.03 | 1.8 |  | 199026.16 | 3 | 1750624.72 | 2.8 | 9870.96 | 3.8 | 7.852 | 6 | 0.335 | 1.8
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7395.06 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7712.15 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1324.95 is outside the allowed range [80.00, 120.00] | Sample |  | 20-90-T |  | 404399.54 | 3 |  | 2422.88 | 1.8 |  | 200186.8 | 2.6 | 1743629.45 | 1.7 | 10114.37 | 1.1 | 44.515 | 3 | 0.929 | 1.9
55 Mn [ He ] : CPS RSD value = 5.63 is over the allowed maximum = 5.0056 Fe [ He ] : CPS RSD value = 5.79 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7357.47 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7868.70 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1282.07 is outside the allowed range [80.00, 120.00] | Sample |  | DI H2O |  | 13376.67 | 5.6 |  | 79 | 8.8 |  | 199169.21 | 2.7 | 1779022.84 | 2 | 9787.03 | 10 | 1.321 | 6.3 | 0.018 | 14.8
56 Fe [ He ] : CPS RSD value = 5.53 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7403.16 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7762.21 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1401.26 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 7108.65 | 3.6 |  | 602.35 | 5.2 |  | 200405.89 | 3.8 | 1754948.15 | 3.2 | 10696.92 | 5 | 0.629 | 4.5 | 0.221 | 5.5
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7481.82 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 7716.10 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1231.85 is outside the allowed range [80.00, 120.00] | Sample |  | 100 ppb check |  | 1051463.16 | 2.1 |  | 273693.26 | 2.2 |  | 202535.27 | 3.4 | 1744522.37 | 3.3 | 9403.71 | 5.5 | 115.993 | 2.2 | 106.301 | 2.2
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7345.47 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 8102.91 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : CPS RSD value = 9.17 is over the allowed maximum = 5.00193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1373.47 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 7865.73 | 3.5 |  | 26000.47 | 2.2 |  | 198844.31 | 3.1 | 1831976.22 | 2.6 | 10484.79 | 9.2 | 0.712 | 4.3 | 10.087 | 2.2
45 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7331.14 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 8208.94 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1293.48 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 4357.47 | 1.5 |  | 9767 | 1.2 |  | 198456.42 | 2.3 | 1855948.25 | 2.1 | 9874.17 | 8.8 | 0.325 | 2.3 | 3.781 | 1.2
56 Fe [ He ] : CPS RSD value = 5.26 is over the allowed maximum = 5.0045 Sc ( ISTD ) [ He ] : ISTD Recovery Percent value = 7408.82 is outside the allowed range [80.00, 120.00]103 Rh ( ISTD ) [ He ] : ISTD Recovery Percent value = 8056.24 is outside the allowed range [80.00, 120.00]193 Ir ( ISTD ) [ He ] : ISTD Recovery Percent value = 1241.47 is outside the allowed range [80.00, 120.00] | Sample |  | blk |  | 2650.33 | 5.3 |  | 5710.81 | 0.8 |  | 200559.28 | 3.3 | 1821425.03 | 4.1 | 9477.12 | 2.8 | 0.136 | 11.4 | 2.206 | 0.8
Slope
y-Axis
LOD
LOQ